Rat Cage Lid Modification Procedure

Parts List:

| Name | Part Number | Vendor / Part Type | Qty / Unit |
| --- | --- | --- | --- |
| Rat Cage Lid |  | Purchased-Techniplast | 1 |
| RC Panel Front | J009859 | Custom machined | 1 |
| RC Panel Side | J009838 | Custom machined | 1 |
| RC Bracket Front | J009600 | Custom sheetmetal-GoProto | 1 |
| RC Bracket Long | J009835 | Custom sheetmetal-GoProto | 1 |
| RC Bracket Side | J009836 | Custom sheetmetal-GoProto | 1 |
| Card Holder Bracket | J010308 | Custom sheetmetal-GoProto | 1 |
| bottom taper section of right water bottle holder |  | Purchased-Part of Techniplast Lid | 1 |
| Rat Water Bottle Transition |  | Custom machined | 1 |
| Rivets - .063” - .125” | 97447A015 | McMaster | 10 |
| Rivets - .126” - .187” | 97447A020 | McMaster | 10 |
| Rivets - .188” - .25” | 97447A025 | McMaster | 2 |
| 18-8 SS No. 4 Washer | 98019A309 | McMaster | 8 |
| 18-8 SS No. 5 Washer | 90945A712 | McMaster | 10 |
| Epoxy, 3M DP125 | 7467A65 | McMaster | 1 |

Tools List:

| Name | Part Number | Vendor | Quantity/Unit |
| --- | --- | --- | --- |
| Angle Grinder/Dremel |  |  |  |
| 1/8” Drill Bit |  |  |  |
| #28 Drill Bit |  |  |  |
| Power Drill |  |  |  |
| Pliers |  |  |  |
| Epoxy Applicator Gun | 74695A71 | McMaster | 1 |
| Masking Tape |  |  | 1 |
| Binder Clips (optional) |  |  | 2 |

Procedure

- Cut Cage Lid

  - Cut water bottle holder sections off the lid – cut lines shown in red on right picture

  - The left and right sides will be cut at different locations

  - Cut right bottle holder section at thick red line in photo below – there will be a natural seam to follow

  - Save the bottom taper section of the holder and metal flap assembly and for reattachment to modified cage

  - Grind off the front card holder on the left side of the cage

- Assemble Flat Cover

Parts:

| Part Number | Quantity |
| --- | --- |
| J009859 | 1 |
| J009838 | 1 |
| J009600 | 1 |
| J009835 | 1 |
| J009836 | 1 |
| 97447A015 | 10 |
| 90945A712 | 2 |

  - Rivet front panel (J009859) and side panel (J009838) together using the long bracket (J009835) and six 97447A015 rivets

    - Proper orientation of the panels, with respect to each other:

    - The bracket is not symmetrical – the side with square corners should be attached to the front panel

    - Bracket should be attached on the inside of the panels, with rivets attached from the outside, so the flat head is on the outside of the plastic

  - Rivet front bracket (J009600) to the front panel with two 97447A015 rivets

    - Bracket should be on the inside of the panel, with the bracket extending down below the panel

    - Rivets attached from the outside, so the flat head is on the outside of the plastic

  - Rivet side bracket (J009836) to the side panel with two 97447A015 rivets

    - Bracket should be on the outside of the panel, with the short side of the “z” attached to the panel and the bracket jutting out away from the panel

    - Rivets will be attached through the bracket into the panel – use two washers (90945A712) behind the panel

- Attach Flat Cover Assembly to Cage

Parts:

| Part Number | Quantity |
| --- | --- |
| 97447A020 | 10 |
| 90945A712 | 8 |
| J010308 | 1 |
| 97447A025 | 2 |

  - Use cover assembly as guides to drill mounting holes through cage lid – hole locations marked in red

    - Flat panel assembly mounts on left side of lid when looking at the front

    - Drill initial holes with 1/8” drill bit

    - Can be helpful to hold the cover assembly in place with binder clips while drilling

  - Remove cover assembly and enlarge holes with #28 drill bit

  - Attach cover assembly to lid

    - The eight rivets going through the front and side walls, as opposed to the bottom rim, should have backing washers behind the lid plastic – locations shown in red

    - The washer for the hole circled in blue will need to be cut into a D-ring to fit in the compressed space along the thin wall

    - The card holder bracket (J010308) should be riveted to the front lip of the lid along with the cover using two 97447A025 rivets

    - At all ten other points of attachment, the 97447A020 rivets are sufficient

    - The two rivets through the bottom lip on the left side of the lid need to be ground flush to the plastic on the inside of the lid

- Assemble Water Bottle Holder Cover and Attach to Cage

Parts:

| Part | Quantity |
| --- | --- |
| rat water bottle transition | 1 |
| bottom taper section of right water bottle holder | 1 |

  - Epoxy bottom taper section of right water bottle holder to the bottom of the machined rat water bottle transition part

  - Epoxy this combined part onto the rest of the cage where the right water bottle holder was originally located

- Epoxy

  - Seal the corners of the covers with epoxy (7467A65)

    - Circled corners need epoxy

Note: There is epoxy all the way around the cut opening in the left side of the lid in these photos – that is leftover from a different modification. Do not replicate this

    - Tape over the openings

    - Inject epoxy with applicator gun from the inside of the lid

    - Stand on end to prevent epoxy from running and let dry 24 hours before removing tape